Revised 2009

PbS/Ag2S Based Ion Selective Electrode Determination of Lead/Hydroxide Equilibria

While lead hydroxide does not give a good quantitative measure of lead, particularly in the presence of competing species it is an important reaction of lead in all analytical determinations.  The manipulation of the test solution for further, instrumental, analysis, nearly always requires a pH adjustment.  If the pH of the solution is too high some lead may precipitate out.  If the instrumental method is sampling only the liquid phase of the system then the precipitate will not be measured and a negative determinate error is imparted to the entire analytical scheme.

SYNOPSIS:  This lab is designed to illustrate the role of pH in detecting the presence of lead.  Students will monitor the lead concentration in solution using a lead ion selective electrode (ISE)  and will simultaneously monitor the pH with addition of standard KOH solution.  Key concepts in this lab will be the effect ionic strength on the measurements, ion selective electrodes,  and calculation of formation constants for lead hydroxide species.  This lab will also give further practice on the use of spreadsheet calculations and manipulations.

READINGS 	 Read pages 273-276 in Critical Reviews on solution soluble chelates and pages 291-293 on ion selective electrodes.  An attached article illustrates the design, construction, and function of a polymer based ISE based on chemistry similar to the dithizone UV-Vis lab.

Testable Skills

1.	Construct alpha plots given the appropriate information.

2.	Create a calibration curve

3.	Calculate selectivity of the method

4.	List/Define limits and advantages of the method

5.	Explain the chemistry which results in a signal

PRE-LAB

Before coming to lab the student should

1.	Calculate an alpha fraction plot vs pH for lead.

2.	Use literature data for Cd in chloride system to calculate sequential formation constants.

Complex equilibria and calculations of formation constants from potentiometric data

A metal ion can bind to a ligand (L=Cl-, CN-, OH-, etc.) in sequential steps.

Reaction	     Constant	Conc in terms of stepwise and overall formation constants

[1]	M + L = ML		K1	[ML] = K1[M][L]

[2]	ML + L = ML2	K2	[ML2] = K2[ML][L] = K1K2[M][L]2 =  $2[M][L]2

[3]	ML2 + L = ML3	K3	[ML3] = K3[ML2][L] = K1K2K3[M][L]3 =  $3[M][L]3

[n]	ML(n-1) + L = MLn	Kn	[MLn]	= Kn[ML(n-1)][L] = K1K2...Kn[M][L]n = $n[M][L]n

Each reaction has been solved for the equilibrium concentration of the appropriate species in terms of the sequential or step-wise equilibrium constants, Ki, and in terms of the overall equilibrium constant, $i.

A mass balance for all forms of the metal is written as:

CM = [M] + [ML] + [ML2] + [ML3] ...+ [MLn] 						(1)

which can be replaced by the concentration terms shown above:

CM = [M] + K1[M][L] + $2[M][L]2 + $3[M][L]3 + ..... $n[M][L]n				(2)

We can define a denominator “D” as:

D = CM/[M] = 1 + K1[L] + $2[L]2 + $3[L]3 + .....$n[L]n					(3)

To calculate the fraction of the total amount of metal in each form we define the “alpha” fractions:

"0 = [M]/CM = 1/D										(4)

"1 = [ML]/CM = K1[M][L]/CM = K1[L]/D							(5)

"2 = [ML2]/CM = $2[M][L]2/CM = $2[L]2/D							(6)

"3 = [ML3]/CM = $3[M][L]3/CM = $3[L]3/D							(7)

"n = [Mln]/CM = $n[M][L]n/CM = $n[L]n/D							(8)

Typically the alpha fractions are plotted against the ligand concentration.  This allows one to see what form of the metal is most prevalent at any given ligand concentration.

Figure 1  Alpha Plot of lead speciation as a function of pH

Figure 1 is a plot of only two the lead species, Pb2+, and Pb(OH)+,  as a function of pH..  You will make a full plot of all alpha fractions using a spreadsheet.

Pb/Hydroxide alpha plot

Using a spreadsheet create and alpha plot fraction for lead using the following constants:

logK1 = 7.82

Log$2 = 10.88

Log$3 = 13.94

Log$4 = 16.30

This is easiest done with the following set up

Column A:	pH values from 1 to 14 in some small increment

Column B:	[OH-] values corresponding to the pH

Column C:	D  = 1 + K1[OH-] + $2[OH-]2 + $3[OH-]3 + .....$n[OH-]n

Column D:	"0 = 1/D

Column E:	"1 = K1[OH-]/D

etc.

Ion Selective Electrodes

In potentiometric methods the selective charge distribution across a membrane is monitored as a potential.  Charge distribution arise due to two processes, interfacial equilibria, and membrane mobility of the ion (see Figure 2).

Figure 2:  Diagram of solid phase and solution phase processes for the lead ion selective electrode.

In Figure 2 there is a mobility of Ag+ ions through a mixed crystal of PbS/Ag2S that is controlled by solubility of the two crystals and by the bulk solution concentration.  If on the interior of the crystal is a fixed solution of silver, and if Ag+ moves across the crystal in response to a concentration gradient, then charge will be differentially displaced due to movement of cations but not of anions.  Since voltage is defined as the amount of charge, Q, stored over the total possible stored charge (capacitance), C:

(9)

a voltage develops which can be measured across the crystal.  The charge stored depends upon the mobility of Ag+ within the crystal.

When all goes well and the silver ion activity is controlled by lead ion activity in solution then the response of the electrode to the lead ion activity is given by the Nikolsky equation:

(10)

where E is the measured voltage (V),  the constant relates to the internal filling solution, R is the natural gas constant, T is the temperature in Kelvin, F is Faraday's constant,  zi is the valence of the ion, i, (here Pb2+) measured, and ai is the activity of the ion measured. Temperature is clearly important:

Assuming room temperature and combining constants we find:

(11)

The activity of the ion is related to the concentration of the ion via the activity coefficient:

(12)

The activity coefficient is a function of the ionic strength, I,  of the solution:

(13)

where the ionic strength of a solution is affected the concentrations of all of the ions and their charge, z.

(14)

Clearly equation 11 indicates that the voltage measurement will be highly dependent upon the ionic strength of the solution, therefore it is often suggested that an ionic strength buffer be used.

Equation 11 also indicates that the electrode may respond directly to other ions which are mobile within the solid state membrane.  This particular type of lead ion selective electrode responds directly to protons which are also mobile within the crystal lattice.   In Figure 3 (from the Orion Lead Ion Selective Electrode manual) it can be observed that at pH 8 it is theoretically possible to measure 10-6 M lead ion, while at pH 2 the limit of detection is approached at ~10-3 M lead ion.

Figure 2=3:  pH dependence of the electrode signal as a function of obth pH and lead ion concentration

GLASSWARE

2 erlenmeyer flasks

EQUIPMENT

1 pH meter

1 voltmeter (for use with the ISE)

1 SCE electrode (for use with the ISE)

Stir box, stir bars

SOLUTIONS

0.1 M NaOH

Ionic strength adjustor

0.01 M NaNO3

Methanol Formaldehyde,

Add three drops of 36% formaldehyde to 1,000 mL reagent grade methanol.  What is the purpose of this reagent?

CAUTIONS AND PROBLEMS

The ion selective electrode responds to free sulfide, to lead, and to pH effects.  As a consequence the pH must not drop below pH 4.5, nor should it exceed pH 8.  When the pH exceeds these limits the calibration curve obtained will not be correct.  The method is further complicated by the slow equilibration time required for the electrode.

PROCEDURE

Create calibration solutions for lead varying between 10-6 and 10-1 M total lead concentration.  One set of solutions should be buffered to pH 2.5 and the second set to pH 6.

A.	Temporal response and calibration curves

1.	Connect leads of the Saturated Calomel Electrode and Pb ISE to voltmeter.  Check which way you connect them and continue to connect in exactly the same manner in any subsequent experiments, otherwise your reading will change from positive to negative.

2.	If Pb electrode is dirty, polish gently on polishing strips.

3.	Rinse electrodes, blot dry.

4.	Measure each of the standard solutions for both pH and mV.  You may have to wait up to or more than ½ hour for the voltage to stabilize.  Monitor the pH and mV at 1 minute intervals.  When the change in mV is less than 0.5 mV between minutes readings you may be near an equilibrium.  The best way to tell if you have come to equilibrium is to plot in lab as you acquire the data the mV reading vs time.  The pH should be identical for each of the standards and should be 4.5 otherwise you will not get a calibration curve when you are done.

5.	You will need to take the mV reading of one of the standards at least three times in order to determine an experimental standard deviation necessary for your LOD calculation.

6.	For your calibration curve take the final, stable mV reading.  Plot the mV reading vs the log[Pb] of the solution.

Calculate the calibration curve from this data.  Be sure to allow for dilution and for the fact that the response is with the logarithm of concentration.

mV = A + Blog[Pb2+]

B.	Construction of an Experimental alpha plot

1.	Begin with a stock solution of lead

2.	Measure both the pH and the mV reading (mV measured between SCE and PbISE) as 10 µL additions of 1.00 M NaOH solution is added.  Stir during the additions.  You should have about 30 points between pH 5 and 8.

3.	Create a table of your experimental data

column A	volume NaOH added

column B	CM	= (initial vol lead)(inital conc.)/total volume

column C	mV

column D	[Pb2+]	= 10(mV - A)/B

column E	pH

4.	From this table determine the experimental alpha fraction, "o,  by dividing the experimental free lead [Pb2+] by the experimental total lead, dilution corrected, Cm.  Plot the experimental alpha as a function of the pH.  Superimpose this plot on your “theoretical plot”.

REPORT

1.	Plot of theoretical alpha plot of lead equilibria.

2.	Does your calibration curve for the lead ISE at pH 6 have the “right” slope?  Why or why not?

3.	Does your calibration curve for the lead ISE at pH 2.5 have the “right” slope? Why or why not?

4.	Why did we use NaClO4 or NaNO3 instead of NaCl for the ionic strength buffer?

5.	Was the assumption of a constant ionic strength valid? (Calculate the ionic strength for each addition of reagents.

6.	Why did we check to make sure that the solutions had a pH<5 for the calibration curve?

7.	What is the purpose of the added methanol/formaldehyde?  Hint: Something about the chemistry of S in air.

8.	Why does there seem to be a time response to the mV readings?

9.	Compare your experimental alpha plot to your theoretical lead vs pH alpha plot.  Do the plots match?  Why or why not?

10.	What implications does this chemistry have for lead analysis?

11.	What implications does this chemistry have for lead in the pipes to your household?